## Slide 1
Real-Time AI ECG Arrhythmia Detector
Complete Project Explanation
A beginner-friendly guide to every aspect of this project
No prior knowledge of AI, medicine, or programming is assumed

## Slide 2
What You'll Learn ( PADHAI LIKHAI KARO 
1. What Problem Does This Solve?
2. What is an ECG?
3. What is an Arrhythmia?
4. The 5 Types of Heartbeats We Detect
5. Where Does the Data Come From?
6. How Data is Prepared (data_prep.py)
7. How the AI Model Works (train_model.py)
8. The Dual-Input Architecture
9. How the Dashboard Works (app.py)
10. Noise Robustness Testing
11. External Validation (INCART)
12. Key Results & What They Mean

## Slide 3
1. What Problem Does This Solve?
Imagine you're a nurse in a hospital ICU. You're watching 8 patient screens simultaneously. Each screen shows a squiggly green line - the patient's heartbeat. Your job is to spot the ONE dangerous heartbeat among thousands of normal ones.
Miss it, and the patient could die.
Hospitals report that nurses ignore up to 90% of alarms due to "alarm fatigue" - too many false alerts.
OUR SOLUTION
An AI system that watches every single heartbeat, classifies it instantly (< 5 milliseconds), and only alerts doctors when something is genuinely wrong.
It never gets tired. Never gets distracted. Catches 98.8% of dangerous heartbeats.

## Slide 4
2. What is an ECG?
An ECG (Electrocardiogram) records the electrical activity of your heart.
Every time your heart beats, an electrical signal travels through it, making the heart muscles contract and pump blood.
An ECG machine captures these signals through small sticky patches (electrodes) on your chest. The result is the familiar 'heartbeat line' on hospital monitors.
The PQRST Wave
Each heartbeat creates a specific pattern:
P wave - Upper chambers contract
QRS complex - Main chambers contract (big spike)
T wave - Heart recovers for next beat
R-peak - The tallest point, marks each beat
Key Numbers
Signal measured in millivolts (mV) - tiny electrical voltages
MIT-BIH records at 360 samples per second (360 Hz) - 360 measurements every second
A 30-minute recording = 648,000 data points per patient
Our AI processes each heartbeat as a window of 250 samples (~0.7 seconds)

## Slide 5
3. What is an Arrhythmia?
An arrhythmia is any heartbeat that doesn't follow the normal rhythm.
Think of your heart like a drummer keeping a steady beat. An arrhythmia is like the drummer suddenly hitting an extra beat, skipping a beat, or beating in a weird pattern.
Some arrhythmias are harmless (everyone has occasional skipped beats). Others can be life-threatening.
The 5 Types We Detect (AAMI Standard)
N - Normal Beat: Heart beating normally from the right place [SAFE]
S - Supraventricular: Beat from the upper chambers instead of the normal pacemaker [MODERATE]
V - Ventricular (PVC): Beat from the lower chambers - the WRONG place [DANGEROUS]
F - Fusion: A normal beat and ventricular beat happen simultaneously [CONCERNING]
Q - Unknown/Paced: Unclassifiable beat, or from an artificial pacemaker [DEPENDS]
PVCs are the most clinically dangerous. If the abnormal electrical source wins too often, it can cause ventricular fibrillation (heart quivers instead of pumping) - fatal within minutes.

## Slide 6
4. Where Does the Data Come From?
MIT-BIH Arrhythmia Database
The gold standard for ECG research since 1980
48 patient recordings, 30 minutes each
Recorded at 360 Hz (360 samples/second)
Every beat manually labeled by 2 cardiologists
Used in 5,000+ research papers
Free from PhysioNet (no licensing issues)
INCART Database (External Test)
Used to prove our model generalizes
75 patients from St. Petersburg, Russia
Different hospital, different equipment
173,476 beats - completely unseen data
Different sampling rate (257 Hz vs 360 Hz)
WHY PATIENT-LEVEL SPLITTING MATTERS
Training: 43 patients | Testing: 5 completely separate patients
The AI has NEVER seen the test patients' hearts during training. This is critical because:
If you mix beats from the same patient in train and test, the AI memorizes the patient's unique heart shape - not the arrhythmia. Many papers do this and report misleading 99%+ accuracy. Our 82% on truly unseen patients is more honest and more useful.

## Slide 7
5. How Data is Prepared (data_prep.py)
This is the FIRST script you run - it downloads and processes the raw ECG recordings
Step 1: Download
Download 48 patient recordings from PhysioNet (free medical data repository)
Step 2: Segment Beats
Find each R-peak (heartbeat marker) and cut a 250-sample window around it: 100 samples before + 150 after. Each window = 1 heartbeat.
Step 3: Normalize
Scale each beat to 0-1 range so the AI focuses on SHAPE, not amplitude. Different patients have different signal strengths.
Step 4: RR Features
Compute 4 timing features per beat: pre_RR (time since last beat), post_RR (time to next), local_avg_RR (average nearby), prematurity ratio.
Step 5: Label
Map cardiologist annotations to 5 AAMI classes: N, S, V, F, Q.
Step 6: Save
Save as .npy files: X_train (97,714 beats), y_train (labels), RR_train (timing features), plus test set and demo recordings.

## Slide 8
6. How the AI Model Works (train_model.py)
What is a Neural Network?
A neural network is like a complex filter. You show it thousands of examples with correct answers, and it slowly adjusts its internal settings until it can predict the right answer on its own.
Like teaching a child dog breeds with photos - after enough examples, they recognize breeds in new photos they've never seen.
What is a 1D-CNN?
CNN = Convolutional Neural Network. A small 'sliding window' moves across the heartbeat signal, computing how much each part matches patterns it's looking for.
Layer 1: Detects simple features (spikes, slopes) Layer 2: Detects QRS patterns, T-wave shapes Layer 3: Detects overall beat type (Normal vs PVC)
Residual (Skip) Connections
Deep networks suffer from 'vanishing gradients' - information gets diluted through many layers.
Skip connections add shortcuts that let information bypass layers. This allows deeper, more powerful networks.
Our model uses 3 residual blocks (64, 128, 256 filters).
Focal Loss
82% of beats are Normal. A lazy AI could predict 'Normal' always and be 82% accurate - but miss every arrhythmia!
Focal Loss forces the AI to focus on rare, hard cases. Easy Normal beats get tiny penalties. Missed PVCs get HUGE penalties.
Like a teacher who spends more time on topics students struggle with.
Data Augmentation
Rare classes have few examples. We create more by slightly modifying existing ones:
1. Add tiny random noise 2. Scale amplitude (+/-15%) 3. Shift timing (+/-5 samples)
Grows training set from 97K to 218K beats.

## Slide 9
7. Dual-Input Architecture - Our Secret Weapon
The Problem
Most ECG AI models look ONLY at the beat waveform shape.
But cardiologists also look at RHYTHM - how the beat relates to its neighbors in time:
'This beat came too early' = might be premature (PVC)
'Long pause after this beat' = compensatory pause (confirms PVC)
Our Solution: Two Inputs
Branch 1 (CNN): Learns what the beat LOOKS like (morphology/shape)
Branch 2 (Dense): Learns what the beat's TIMING tells us (rhythm)
Together, they make better decisions - just like a cardiologist who considers both shape AND rhythm.
Architecture Diagram
INPUT 1: Beat Waveform (250 samples)
Conv1D(64,7) -> ResBlock(64) -> ResBlock(128) -> ResBlock(256) -> GlobalAvgPool
|
CONCATENATE -> Dense(256) -> Dropout -> Dense(128) -> Softmax(5)
|
INPUT 2: RR Features (4 numbers)
Dense(16) -> BatchNorm -> Dense(8) ----------------------------------------+
RR features are patient-independent: PVC timing (early beat + pause) is universal regardless of waveform shape.

## Slide 10
8. How the Dashboard Works (app.py)
Tab 1: Live Monitor
Simulates a real hospital bedside monitor:
1. Select a patient recording from dropdown
2. Click Play - ECG scrolls across screen
3. Each beat classified in real-time:
Green = Normal | Red = PVC | Orange = SVE
4. Live heart rate + arrhythmia count
5. Alert log for every abnormal beat
Tab 2: Model Performance
Shows the AI's test results:
Confusion Matrix - predictions vs reality grid
Per-class precision, recall, F1-score
Training accuracy curves over epochs
Class distribution charts
How Inference Works (Behind the Scenes)
1. Load recording -> 2. Feed ALL beats + RR features through model at once (batch prediction) -> 3. Model outputs probabilities [P(N), P(S), P(V), P(F), P(Q)] -> 4. Cache results -> 5. During playback, just display cached results frame by frame
KEY INSIGHT: All inference runs ONCE when you select a recording, not during playback. That's why animation is smooth.

## Slide 11
9. Noise Robustness Testing
Why Test With Noise?
In real hospitals, ECG signals are never perfectly clean. Patients move, electrodes shift, machines interfere. A model that works on clean data but fails on noisy data is useless.
3 Types of Noise We Test
Gaussian - Random static (like radio interference)
Baseline Wander - Slow drift from breathing/movement
Muscle Artifact - High-freq noise from fidgeting
Results: Accuracy at Different Noise Levels
SNR (dB) Gaussian Baseline Wander Muscle Artifact Combined
Clean 81.1% 81.1% 81.1% 81.1%
30 dB 79.3% 81.3% 80.3% 80.7% (typical hospital)
20 dB 15.4% 81.6% 26.3% 25.1%
10 dB 2.1% 81.4% 2.0% 3.1%
0 dB 1.3% 71.3% 9.7% 2.3% (extreme noise)
Key finding: Model is very robust to baseline wander (71.3% even at 0dB!) - the most common real-world noise source.

## Slide 12
10. External Validation (INCART Database)
Why External Validation?
Training and testing on the same database is like studying for an exam using the exact textbook the exam comes from.
We tested on INCART - a completely different hospital (Russia vs USA), different equipment, different patients. This proves real generalization.
INCART Results
87% overall accuracy on 173,476 beats
Normal beats: correctly identified across databases
Arrhythmia detection: limited (expected)
Shows the path forward: domain adaptation needed for multi-site deployment
This result demonstrates scientific rigor. Most hackathon/research projects never test on external data. Showing this - even with imperfect results - proves we understand real-world deployment challenges and sets our project apart from those that only show inflated numbers on their training dataset.

## Slide 13
11. Key Results & What They Mean
98.8%
PVC Recall
Catches nearly every dangerous beat
82.0%
Test Accuracy
On patients NEVER seen in training
608K
Parameters
Lightweight model, runs anywhere
<5ms
Inference Time
Real-time on CPU, no GPU needed
Precision vs Recall (Explained Simply)
RECALL = Of all actual PVCs, what % did we catch?
98.8% = we only miss 1.2% of dangerous beats
PRECISION = Of our PVC alerts, what % were real?
41.5% = some false alarms, but safe
Think SMOKE DETECTOR: Better to have false alarms than miss a real fire. We prioritize HIGH RECALL for patient safety.
Model Evolution
v1: Basic CNN
67.3% accuracy | 96.8% V recall
v2: + Class Weights
81.4% accuracy | 96.8% V recall
v3: + Residual CNN + Focal Loss
81.1% accuracy | 97.2% V recall
v4: + Dual-Input + RR Features
82.0% accuracy | 98.8% V recall

## Slide 14
12. How All Files Connect Together
Pipeline Flow
STEP 1: data_prep.py STEP 2: train_model.py STEP 3: app.py
Download from PhysioNet Load .npy files Load trained model
Segment beats Oversample + Augment Live ECG playback
Normalize + RR features Build Dual-Input CNN Real-time classification
Save .npy files ---> Train with Focal Loss ---> Display alerts
Save .keras model Performance tab
Main Scripts
data_prep.py - Downloads & processes data (run FIRST)
train_model.py - Trains the AI model (run SECOND)
app.py - Dashboard for live demo (run ANYTIME)
noise_robustness.py - Tests noise resilience
external_validation.py - Tests on INCART database
create_ppt.py - Generates pitch deck
run_dashboard.bat - One-click launcher
Output Files (output/ folder)
ecg_model.keras - Trained AI model (2.3 MB)
X_train.npy, y_train.npy - Training data
RR_train.npy, RR_test.npy - Timing features
metrics.json - Accuracy numbers
confusion_matrix.png - Prediction grid
noise_robustness_curve.png - Noise plot
demo_recordings/*.npz - Dashboard data

## Slide 15
Thank You!
Real-Time AI ECG Arrhythmia Detector
Now you understand every part of this project!
Built with Python | TensorFlow | Streamlit | MIT-BIH Database